## Annex 1.A
Supply of medical equipment and consumables for Ambulances
Annex 1.A - Technical Specifications for Medical Equipment and Tools
# | Item | Specification | Supplier Specification | Brand | Manufacturer | Delivery Period (days)
1 | Aneroid Sphygmomanometer - Adult | Components: Delivered in a vinyl or leatherette zipper case Parts of sphygmomanometer: - Rubber bulb: to inflate air bladder - Aneroid pressure gauge (no stop pin) with air pressure- release valve ; easy-to-read, maximum error throughout range ±3mm Hg ; Connector to rubber tube, on the right side below the manometer - Connecting tube: 60cm - Cuff: non-deformable material, washable at 30 °C Velcro fastening enabling tight adjustment around the arm - Inflatable bladder, washable Material: - Tempered copper- beryllium: diaphragm - Stainless steel: handle - Nickel- plated: tube connector and valve - Nylon: cuff +Velcro - Latex: bulb, bladder Specifications: Instrument for measuring arterial blood- pressure by means of a manometer and an inflatable cuff Adult size Cuff wide: length 57cm, width 14cm Inflatable bladder: length 22cm, width 10cm Graduated scale Ø65mmHg up to 300mmHg Paediatric size 3 interchangeable cuff size: . New born: cuff length 15cm, width 5cm . Infant: cuff length 23cm, width 7.2cm . Child: cuff length 35.5cm, width 10cm Graduated scale Ø40 mmHg up to 300mmHg Packaging: Unit presentation, in a nylon holdall The following information must appear on the packaging: - designation of item - name and address of manufacturer - manufacturer's certificate of guarantee and accompanying instructions for use enclosed inside packaging Warranty : Two years or more warranty
2 | Stethoscope | 1. Guarantee: 2 years or more 2. Acoustic tube: made in PVC; flexible; soft 3. Eartips: elastic 4. Chest piece: made of machined stainless steel; Diaphragm in fibreglass
3 | oximeter fingertip pulse | 1. For use in Health Care facility 2. Measurement range: SpO2 0 – 100%; pulse 30 – 250 beats/minute 3. Accuracy: SpO2 70 – 100%, ±2% / pulse 30 – 250 bpm, ±2 bpm 4. SpO2 sensor: Adult SpO2 sensor with cable- two nos per monitor and Pediatric SpO2 sensors- one no. per monitor 5. Operating conditions: 10°C to 55°C, ≤75% RH, 700 – 1060 hPa ambient pressure 6. Storage and transport conditions: - 40°C to 60°C; ≤95% Relative Humidity; 500 – 1060 hPa ambient pressure 7. Display reports alarms, system errors, sensor failure/disconnection and battery status Auto power-off function 8. Conform with ISO 80601-2-61:2011 or EN 60-601-1-2 in terms of Electrical, Mechanical, Environmental safety and Electromagnetic compatibility, as well as labeling 9. Conform with ISO-10993 in terms of biocompatibility 10. Vibration testing: IEC 68-2-6/68-2-3I 11. Warranty: 2 year or more
4 | Glucometer | Display: LCD with backlight. Sample size: 0.6 ml. Power supply: Memory capacity: 720 diary records plus 32 control tests. Test averages: Dimensions: 3.2" x 1.9" x 0.8" Weight: Approximately 1.4 oz with batteries. Data transfer: Bluetooth® wireless technology and micro USB.
5 | Glucometer strips | High quality Should be of the same type of the glucometer supplied
6 | SAM Splint | Aluminum foam splint (SAM type), malleable and lightweight, with aluminum core (~0.4 mm) and closed-cell foam coating; waterproof, non-toxic, radiolucent, reusable; easily shapeable and rigid when curved; suitable for fracture immobilization; standard size approx. 90–92 cm × 10–11 cm; can be cut to size
7 | Rigid cervical collar | Rigid cervical collar (adult/pediatric), adjustable multi-size, made of radiolucent plastic with foam padding; lightweight, non-toxic, X-ray compatible; with chin support, tracheal opening, ventilation holes and Velcro closure; designed for pre-hospital/ambulance use, providing rapid application and effective cervical spine immobilization during emergency transport
8 | Ambulance Patient Trolley Bed with adjustable base. Best quality in the local market | As Per Attached Annex 1.A.1
| Quality Specifications: | Please refere to the Quality Specifications section in the Request for Qoutation file.
Delivery Period (days):
Company Name:
Contact Person:
Phone Number:
Date:
Signature and Stamp: